Current consumption and associated voltage drop at the remote end of 120 feet of CAT5E solid copper conductor shielded cable, used to extend power and I2C bus communications to a remote PNI RM3100 magnetometer.

Raspberry Pi 3B+ host, Buster OS, KD0EAG simplei2c.c, with SparkFun Pi Quiic Hat, Pi 9615 differential I2C bus extenders, PNI 3100 test board on KD0EAG simple adapter board.

Tektronix DC-50MHz current probe and Tektronix DPO2024 oscilloscope.
04-26-2020 J. Madey

Ch1: 5mA / div, CH2: VDA/B voltage response to remote current transients during I2C data transaction; DC coupled 50mV/div using in-line offset voltage to maintain trace on screen. Kemet 220uf 6.3V Tantalum Polymer SMD cap, 70mohm ESR, across VDA/B to ground on PC board. Voltage 'ripple' just at borderline 50mV pp spec of 3100 board.

Same as above but without 220uF VDA/B bypass cap on remote PC board; traces rearranged slightly for better viewing.
Current spikes peak to ~28mA, representing the active load of the 9615 differential driver at the remote end, reduced to ~20mA by storage capacity of the 220uF bypass. VDA/B Voltage noise 'ripple' now in excess of 70-80mV pp. MLCC bypass caps on SF board and 3100+adapter board provide some suppression.
Voltage transients consistent with ~3.3ohm resistance of the 120 foot paralleled #24 twisted pair power/ground circuit. Would generally recommend remote end LDO regulator for 9615 extender designs plus sufficient low ESR cap bypassing at remote end to reduce high frequency noise on long cable runs.